## Slide 1: The OpenPET Firmware and Software Development
Qiyu Peng

## Slide 2: Guidelines for firmware/software design
Stable and reliable
Flexible, compatible and scalable
Simple and easy to use
Upgradable
Well-documented
Readable and maintainable

## Slide 3: System-level hardware structure

## Slide 4
System-level firmware and software structure

## Slide 5: Programming Tools / Environment
 | DB | DUC | MB | CUC | Host PC
FPGA Firmware | X | X | X | X | 
Embedded MCU Software (NIOS) |  | X | X | X | 
PC Software |  |  |  |  | X
X: VHDL (Altera Quartus II ver. 12.0)
X: C (Altera Nios II EDS ver. 12.0)
X: C (NI LabWindows CVI 2009)

## Slide 6: C programing style guidelines
Based on “C Style Guide” suggested by Software Engineering Laboratory (SEL), the National Aeronautics and Space Administration (NASA) in 1994.
Rules and guidelines are used to guarantee the readability and maintainability of the C source codes (details are in my word file)

## Slide 7: C programing style guidelines
Doxygen JavaDoc style documentation generated from “OpenPETexampleCStyle.c”.

## Slide 8: C programing style guidelines
Doxygen JavaDoc style documentation generated from “OpenPETexampleCStyle.c”.

## Slide 9: VHDL programing style guildlines
12 rules and 14 guidelines are used to guarantee the readability and maintainability of the VHDL source codes (details are in my word file)
OpenPET VHDL program Style template for file prolog
Rules and guidelines are developed from Coding Style Guidelines suggested by Xilinx and Altera

## Slide 10
OpenPET VHDL program Style template for file prolog, Modular design

## Slide 11
OpenPET VHDL program Style template for file prolog, timing diagram of input and output
http://users.eecs.northwestern.edu/~seda/coding_guidelines_013003.pdf
http://www.cs.tut.fi/~ege/Misc/dcs_vhdl_coding_rules_es_v4_4.pdf
http://www.altera.com/literature/hb/qts/qts_qii51007.pdf

## Slide 12
Host PC software
Modular Design
Simple and easy to use
Upgradable
Main Functions
System configuration
System diagnosis
Data acquisition
Data analysis

## Slide 13: Graphical User Interface (GUI) design
Top level design

## Slide 14: Graphical User Interface (GUI) design
Analog gain settings
CFD settings
Energy window settings
Serial communication
Energy data analysis
TDC calibration
Data acquisition
Data analysis SW package
(currently in Matlab)
Crystal decoding
Delay compensation
Time analysis
Sinogram analysis
System configuration database management
Flood map display
(for block detector)
SSPM event display
Chinh Vu and Michael Malus also contributed to the GUI design

## Slide 15: Graphical User Interface (GUI) design
System characterization

## Slide 16
Nios software
Modular design
Simple and easy to use
Upgradable
Main Functions
System configuration
System monitoring
System diagnosis

## Slide 17: Firmware function and framework
CUC/MB/DUC/CDUC firmware
DB firmware

## Slide 18: Event words and status words
In general, there are two kinds of words in the OpenPET list mode data:
Event words, include
-- Single event words, and
-- Coincidence
Status words transmit information other than the event data to the host PC, such as:
-- Event data format
-- Time information
-- Single or coincidence event rate
-- System configuration and status
-- Application-specific information such as patient bed position, transmission source position, system temperature, ECG signal, gate signal and etc.

## Slide 19: Event words
Event word path
DUC
Generate original single event data
DB
Add 3-bit DB
address to the
single event data
MB
Add 3-bit DUC
address to the
single event data
CUC
Add 3-bit MB address to single event data.
Generate coincidence event data (optional).
Insert status words to the data stream.
Host PC
Generate original single event data
DB
CDUC
Add 3-bit DB address to single event data.
Generate coincidence event data (optional).
Insert status words to the data stream.
Host PC
DUC
Generate original single event data
DB
Add 3-bit DB
address to the
single event data
CUC
Add 3-bit DUC address to single event data.
Generate coincidence event data (optional).
Insert status words to the data stream.
Host PC

## Slide 20: Event data length
Data from DBDUC, DUC MB and MBCUC
Single event data only
32-bit mode or 64-bit mode.
Data from CUC  host PC
Single event data, coincidence event data and status words.
32-bit mode, 64-bit mode, and 128-bit mode.

## Slide 21: Data paths and data mode combinations
DB
DUC
MB
CUC
Host PC
32 bits
32 bits
32 bits
32 bits
DB
DUC
MB
CUC
Host PC
32 bits
32 bits
32 bits
64 bits
DB
DUC
MB
CUC
Host PC
64 bits
64 bits
64 bits
64 bits
DB
DUC
MB
CUC
Host PC
64 bits
64 bits
64 bits
128bits
16 system clocks
8 system clocks
Oscilloscope mode

## Slide 22: DBDUC, DUC MB and MBCUC
32-bit data (8 system clocks)
General definition
31
30
29
28
27
26
25
24
23
22
21
20
19
18
17
16
15
14
13
12
11
10
9
8
7
6
5
4
3
2
1
0
MB address
DUC address
DB address
Block address
Crystal ID
Payload
1
Crystal ID

## Slide 23: DBDUC, DUC MB and MBCUC
32-bit mode (8 system clocks)
Payload definition
-- Time mode – Bit 10~0, TDC data bits (11 bits, LSB: 50ps).
-- Energy mode – Bit 10~0, energy data bits (11 bits).
-- Test mode1 – Bit 10~0, test pattern data: 10101010101.
-- Test mode2 – Bit 10~0, test pattern data: 01010101010.
-- Anger-logic flood map plot mode – Bit 18, not used,
Bit 17~9, Y value (9 bits), Bit 8~0, X value (9 bits).
-- User-defined mode 1 - 11 bits of payload
-- User-defined mode 2 - 14 bits of payload
-- User-defined mode 3 - 17 bits of payload
-- User-defined mode 4 - 22 bits of payload
-- User-defined mode 5 - 25 bits of payload
-- User-defined mode 6 - 28 bits of payload
Standard
User-defined

## Slide 24: DBDUC, DUC MB and MBCUC
32-bit mode (Oscilloscope mode)
General definition
31
30
29
28
27
26
25
24
23
22
21
20
19
18
17
16
15
14
13
12
11
10
9
8
7
6
5
4
3
2
1
0
MB address
DUC address
DB address
Channel address
Raw ADC data
1
Data counter
Unused
The maximum length of the raw data train is 64

## Slide 25: DBDUC, DUC MB and MBCUC
64-bit data (16 system clocks)
General definition

## Slide 26: DBDUC, DUC MB and MBCUC
32-bit mode (8 system clocks)
Payload definition
-- Time mode – Bit 10~0, TDC data bits (11 bits, LSB: 50ps).
-- Energy mode – Bit 10~0, energy data bits (11 bits).
-- Test mode1 – Bit 10~0, test pattern data: 10101010101.
-- Test mode2 – Bit 10~0, test pattern data: 01010101010.
-- Anger-logic flood map plot mode – Bit 18, not used,
Bit 17~9, Y value (9 bits), Bit 8~0, X value (9 bits).
-- User-defined mode 1 - 11 bits of payload
-- User-defined mode 2 - 14 bits of payload
-- User-defined mode 3 - 17 bits of payload
-- User-defined mode 4 - 22 bits of payload
-- User-defined mode 5 - 25 bits of payload
-- User-defined mode 6 - 28 bits of payload
Standard
User-defined

## Slide 27: CUC  host PC
Single event mode
In single event data mode (including the 32-bit raw ADC data mode), the formats of the data from CUC/CDUC host PC are the exactly the same that those defined in previous section (DBDUC, DUC MB and MBCUC).

## Slide 28: CUC  host PC
Coincidence event mode
Two 32-bit single event data  one 64-bit coincidence event data
Two 64-bit single event data  one 128-bit coincidence event data
Two 32-bit single event data  one 32-bit coincidence event data
Two 64-bit single event data  one 64-bit coincidence event data
MB address
DUC address
DB address
Block address
Crystal ID
Payload
1
Crystal ID
31
30
29
28
27
26
25
24
23
22
21
20
19
18
17
16
15
14
13
12
11
10
9
8
7
6
5
4
3
2
1
0
MB address
DUC address
DB address
Block address
Crystal ID
Payload
Valid
Crystal ID
Single Event 1
Single Event 2
47
46
45
44
43
42
41
40
39
38
37
36
35
34
33
32
63
62
61
60
59
58
57
56
55
54
53
52
51
50
49
48

## Slide 29: Status word (fixed to 64 bits)
General definition

## Slide 31
64 bits data format
Bit63=0, coincidence data
Bit63=1, not coincidence data
Bit62=0, default data format
Bit62=1, User-defined data format
Bit62=1, single event data format
Bit62=0, status data format
Bit61~58=0000, time mode data
Bit61~58=0001, energy mode data
Bit61~58=0010, raw ADC data
Bit61~58=0011, Anger logic data
Bit61~58=0100, Test mode1 data
Bit61~58=0101, Test mode2 data
Bit61~58=0110~1010, reserved
Bit61~58=1011~1111, user-defined
Bit61~57=00000, time status word
Bit61~57=00001, event rate status word
Bit61~57=00010, temperature status
Bit61~57=00011~10000,reserved
Bit61~57=10001~11111, user-defined
Branches in 64-bit list mode data file

## Slide 32
Start
Read 8 bytes of data
Check the two MSB of the 8 bytes of data
Find pattern “00”, …, “01”?
An event format status word (00000)?
Yes
Yes
Determine the event data format
No
End
No
Algorithm to decode list mode data file